## Sheet1
Points | Description | Existing in our current software | Remarks | Possibility | Doubts | No of days
1 | Analyzer-semi automated bench top device using wet chemistry reagents | TBD
2 | Analyzer should have ability to use both external cuvettes and integrated flow cell | TRUE
3 | Analyzer should have direct test access keys on the key board for routine chemistry parameters. | TBD | Extra buttons not available. Can be done from external keyboard if needed. | Onscreen buttons can be provided for this. With lot of changes.
4 | Analyzer should have more than 190 programmable channels | TRUE
5 | Analyzer must have key board with water proof membrane | FALSE | We donot have keyboard.
6 | Analyzer must have following assay types | TRUE
7 | Analyzer must have following calibration types: Linear, Two-point, Kfactor, Logit-log | TRUE
8 | In kinetic assays, measurement interval should be 1 second | TRUE
9 | Analyzer must have storage for three different calibration for each chemistry | Possible | TBD | To show last 3 calibrations? Give option to user to select calibration.
10 | Three level controls (QC) with day to day Levey Jennings chart stored and displayed. | TRUE
11 | The flow cell must be Quartz | TRUE
12 | The flow cell must have an optical path of 10 mm | TRUE
13 | The flow cell volume should be less than 20 ul | Possible | Current flowcell volume is 32ul. | Performance testing will be needed. Count difference to be checked.
14 | Measurement temperature range should be from 20-40 degree C with variable 1 degree C increment. | TBD | Currently we have 4 set temperatures. 25,30,37,OFF.
15 | Analyzer must have following wavelengths as standard.• 340 rim, 415 mm, 510 rim • 546 nm, 578 nm, 600 mm • 660 cm, 700 rim. | Possible | Current wavelengths: 340, 405, 505, 546, 578, 620, 670 | Option to add 3 filters.
16 | Analyzer should have absorbance range from 0.00-3.0 Abs units. | TRUE
17 | Analyzer resolution must be 0.0001 Abs | TRUE
18 | Analyzer detector should be more than 12 bit silicon photo diode | TRUE | It is 20 bit.
19 | Analyzer must store 1000 results | TRUE
20 | Analyzer must store reaction graphs for previous 10 samples | Possible | We currently store only the last sample. | Do we need to print as well?.
21 | One year QC data can be stored | TRUE
22 | Internal thermal printer should be available. | TRUE
23 | Analyzer should be capable to do multiple testing up to 3 replicates. Should display mean. SD, CV. | TBD
24 | Measuring time programmable from 2-998 seconds for kinetic & two point type tests and delay from 0-999 seconds | TRUE
| Analyzer should be able to perform HbAlc testing | TRUE |  | TBD
| Analyzer should have semi-automatic aspiration of reaction mixture directly into flow cell using peristaltic pump. | TRUE
| Firm will supply one UPS (1 KVA online), Micropipettes-4 (10 p1, 25 pl 500 ql, 1000 Jil), Tip-400 (100 for each pipettes) and reagent Glucose, Urea, Creatinine, SGPT, SGOT, Bilirubin and QC. (1 set of reagent) free of cost at the time installation of equipment | TRUE
| The system should be quoted with standard three year guarantee with all accessories. | TRUE
| CMC CMC shall be given @ 4% of net rate (excluding GST as applicable) and yearly escalation of 5% on last year's CMC price. The CMC may be awarded for five years (on yearly basis) after Guarantee period of three years. | TBD
| Demonstration of functionality should be arrange at Jaipur | TBD
| Quoted model should be USFDA/CE EUROPEAN (Notified Body) approved, certificate must be submitted in technical bid. | TBD | CE Certifiation done.
| Copy of valid CDSCO licence for quoted model must be submitted in technical bid | TBD