## Sheet1
Points | Description | Internal Discussion | Remarks | Possibility | Doubts | No of days
1 | Analyzer-semi automated bench top device using wet chemistry reagents | TRUE
2 | Analyzer should have ability to use both external cuvettes and integrated flow cell | TRUE
3 | Analyzer should have direct test access keys on the key board for routine chemistry parameters. | Possible |  | Can be done by opening Run menu directly for these tender machines.
4 | Analyzer should have more than 190 programmable channels | TRUE
5 | Analyzer must have key board with water proof membrane | TRUE | We donot have keyboard. Our display is waterproof
6 | Analyzer must have following assay types | TRUE
7 | Analyzer must have following calibration types: Linear, Two-point, Kfactor, Logit-log | TRUE
8 | In kinetic assays, measurement interval should be 1 second | TRUE
9 | Analyzer must have storage for three different calibration for each chemistry | Possible | To be introduced for normal instruments as well | Store 3 last calibration. Give user option to select which calibration to be used. | To show last 3 calibrations? Give option to user to select calibration.
10 | Three level controls (QC) with day to day Levey Jennings chart stored and displayed. | TRUE
11 | The flow cell must be Quartz | TRUE
12 | The flow cell must have an optical path of 10 mm | TRUE
13 | The flow cell volume should be less than 20 ul | Possible | Current flowcell volume is 32ul. | Performance testing will be needed. Count difference to be checked. To be tested for normal instruments as wel
14 | Measurement temperature range should be from 20-40 degree C with variable 1 degree C increment. | TRUE | Currently we have 4 set temperatures. 25,30,37,OFF. | No need for these variable temps. Current temperature option are good enough
15 | Analyzer must have following wavelengths as standard.• 340 rim, 415 mm, 510 rim • 546 nm, 578 nm, 600 mm • 660 cm, 700 rim. | Possible | Current wavelengths: 340, 405, 505, 546, 578, 620, 670 | We can add 700 filter at free position.
16 | Analyzer should have absorbance range from 0.00-3.0 Abs units. | TRUE
17 | Analyzer resolution must be 0.0001 Abs | TRUE
18 | Analyzer detector should be more than 12 bit silicon photo diode | TRUE | It is 20 bit.
19 | Analyzer must store 1000 results | TRUE
20 | Analyzer must store reaction graphs for previous 10 samples | Possible | We currently store only the last sample. | Do we need to print as well? No Only display is needed. This feature is only tender specific.
21 | One year QC data can be stored | TRUE
22 | Internal thermal printer should be available. | TRUE
23 | Analyzer should be capable to do multiple testing up to 3 replicates. Should display mean. SD, CV. | Possible | Tender Specific | In result search, user can select the samples for mean, SD, CV,
24 | Measuring time programmable from 2-998 seconds for kinetic & two point type tests and delay from 0-999 seconds | TRUE
25 | Analyzer should be able to perform HbAlc testing | TRUE |  | TBD
26 | Analyzer should have semi-automatic aspiration of reaction mixture directly into flow cell using peristaltic pump. | TRUE